## Slide 1
Mass spectrometry (Test)
Mass spectrometry (MS) is an analytical technique that measures masses of particles and for determining the elemental composition of a sample or molecule.
In a typical MS procedure:
A sample is loaded onto the MS instrument and undergoes vaporization.
2. The components of the sample are ionized by one of a variety of methods (e.g., by impacting them with an electron beam), which results in the formation of charged particles (ions).
3. The ions are separated in an analyzer by electromagnetic fields.
4.The ions are detected, usually by a quantitative method.
5.The ion signal is processed into mass spectra.

## Slide 2
Introduction to Mass Spectrometry (Test)
Sample
introduction
Ionization
Minimize collisions, interferences
Separate
masses
Count ions
Collect results

## Slide 3
Sample Introduction Systems (aka “front ends”)
1) Gas source (lighter elements)
2) Solid source (heavier elements)
3) Inductively coupled plasma (all elements, Li to U)

## Slide 4
Ionization occurs in the ‘source’
Electron Ionization
Gas stream passes through beam of e-,
positive ions generated.
Thermal Ionization
Plasma: Gas stream passes through plasma
maintained by RF current and Ar.
Themal: Filament heated to ~1500C

## Slide 5
Mass Analyzers - the quadrupole vs. magnetic sector
Quadrupole:
Changes DC and RF
voltages to isolate
a given m/z ion.
PRO: cheap, fast, easy
Magnetic Sector:
Changes B and V to focus
a given m/z into detector.
PRO: turn in geometry means
less ‘dark noise’,
higher precision,

## Slide 6
56Fe
very low concentrations
in environmental samples,
but high interest (why?)
Unfortunately, 56Fe has the
same atomic wt as ArO
(40Ar+16O)
Quadrupole measurement =
INTERFERENCE!
Low vs. High – resolution ICPMS and Interferences
HR-ICPMS measurement =
can distinguish 56Fe from ArO
NOTE: most elements can be
distinguished with a low
resolution quadrupole

## Slide 7
Ion microprobe
(or
Secondary
Ion
Mass
Spectrometry
 SIMS)
use an ion beam (usually Cs+1) to “sputter” a sample surface; secondary ions fed into mass spec
20μm

## Slide 8
Accelerator Mass Spectrometry
The AMS at University of Arizona (3MV)
prior to AMS samples were 14C-dated by counting the number of decays
- required large samples and long analysis times
-1977: Nelson et al. and Bennett et al. publish papers in Science demonstrating
the utility of attaching an accelerator to a conventional mass spectrometer
The AMS at LLNL (10MV)
Principle:
You cannot quantitatively remove interferring
ions to look for one 14C atom among several
quadrillion C atoms.
Instead, you
destroy molecular ions (foil or gas)
filter by the energy of the ions (detector)
to separate the needle in the haystack.

## Slide 9
1) Dark Noise - detector will register signal even without an ion beam
- no vacuum is perfect
and
- no detector is perfect
- must measure prior to run to get “instrument blank” if needed
2) Detector “gain” - what is the relationship between the electronic signal recorded
by the detector and the number of ions that it has counted?
- usually close to 1 after factory calibration
- changes as detector “ages”
- must quantify with standards
Cardinal rule of mass spectrometry:
Your measurements are only as good as your STANDARDS!
Standards (both concentration and isotopic) can be purchased from NIST
Hurdles in mass spectrometry (cont.)